## Slide 1: University Astronomy
Professor Don Figer
Instruments and Detectors II

## Slide 2: Lecture Outline
Photon Detection in PN Junctions
  review semiconductors
  PN Junction
  charge collection in PN junctions
Review of CCDs
  definition
  design
  operation
  performance
Review of Infrared Detectors

## Slide 3: pn junctions

## Slide 4: Semiconductors
A conductor has free (unbound) electrons that can flow in the presence of an electric field.
An insulator impedes the flow of electrons.
A semiconductor becomes a conductor if the electrons are excited to high enough energies, otherwise it is an insulator.
  allows for a “switch” which can be on or off
  allows for photo-sensitive circuits (photon absorption adds energy to electron)
Minimum energy to elevate an electron into the conduction band is the “band gap energy”

## Slide 5: Periodic Table
Semiconductors occupy column IV of the Periodic Table
Outer shells have four empty valence states
An outer shell electron can leave the shell if it absorbs enough energy

## Slide 6: Simplified Silicon Band Diagram
Conduction band
Valence band
Eg bandgap
um

## Slide 7: Semiconductor Dopants

## Slide 8: PN Junctions
In a PN junction, positively charged holes diffuse into the n-type material. Likewise, negatively charged electrons diffuse in the the p-type material.
This process is halted by the resulting E-field.
The affected volume is known as a “depletion region”.
The charge distribution in the depletion region is electrically equivalent to a 2-plate capacitor.

## Slide 9: Photon Detection in PN Junctions
A photon can interact with the semiconductor to create an electron-hole pair.
The electron will be drawn to the most positively charged zone in the PN junction, located in the depletion region in the n-type material.
Likewise, the positively charged hole will seek the most negatively charged region.
Each photon thus removes one unit of charge from the capacitor. This is how photons are detected in both CCDs and most IR arrays.

## Slide 10: The Band Gap Determines the Red Limit

## Slide 11: ccds

## Slide 12: CCD Definition
A CCD is a two-dimensional array of metal-oxide-semiconductor (MOS) capacitors.
The charges are stored in the depletion region of the MOS capacitors.
Charges are moved in the CCD circuit by manipulating the voltages on the gates of the capacitors so as to allow the charge to spill from one capacitor to the next (thus the name “charge-coupled” device).
An amplifier provides an output voltage that can be processed.
The CCD is a serial device where charge packets are read one at a time.

## Slide 13: CCD Pixels
Charge motion
Charge motion
Serial (horizontal) register
Parallel (vertical) registers
Pixel
Image area
(exposed to light)
Output amplifier
masked area
(not exposed to light)

## Slide 14: What is This?
Notes: 14

## Slide 15: First Astronomical CCD image
1974 on an 8” telescope
Notes: 15

## Slide 16: CCD in a Dual-Inline Package
Notes: Delta-Doped Charged Coupled Devices (CCD) for Ultra-Violet and Visible Detection
CCDs allow scientists to study one of the least explored windows of the electromagnetic spectrum - the extreme ultraviolet. Until recently, scientists believed there was little point to exploring this spectral region. They thought that the mixture of hydrogen gas and other less abundant gases, which fill the space between stars and is commonly called the "interstellar medium," would absorb virtually all extreme ultraviolet radiation before it became detectable from Earth. Consequently, this region became known as the "unobservable ultraviolet."
This is where stable ultraviolet CCDs step in for crucial space and ground-based astronomy to detect younger, hotter objects. Because of the high quantum efficiency of these devices, very faint objects can potentially be observed. Stability of the device makes it possible to gather reliable data in space-based astronomy.
Scientists will learn much more about hot white dwarf stars - extremely dense stars that represent a final stage of stellar evolution - and young massive stars that are characterized by outflowing, shock-heated winds. Moreover, they will learn about cataclysmic variable stars, binary star systems in which the mass of one star is transferred to the other, causing dramatic changes in extreme ultraviolet brightness. They may even have a chance to probe the enigmatic cores of distant galaxies."

## Slide 17: MOS Capacitor Geometry
A Metal-Oxide-Semiconductor (MOS) capacitor has a potential difference between two metal plates separated by an insulator.
Notes: 17

## Slide 18: CCD Clocking

## Slide 19
pixel
boundary
Charge packet
p-type silicon
n-type silicon
SiO2 Insulating layer
Electrode Structure
pixel
boundary
incoming
photons
Photons entering the CCD create electron-hole pairs. The electrons are then attracted towards the most positive potential in the device where they create ‘charge packets’. Each packet corresponds to one pixel

## Slide 20
1
2
3
+5V
0V
-5V
+5V
0V
-5V
+5V
0V
-5V
Time-slice shown in diagram
1
2
3

## Slide 21
1
2
3
+5V
0V
-5V
+5V
0V
-5V
+5V
0V
-5V
1
2
3

## Slide 22
1
2
3
+5V
0V
-5V
+5V
0V
-5V
+5V
0V
-5V
1
2
3

## Slide 23
1
2
3
+5V
0V
-5V
+5V
0V
-5V
+5V
0V
-5V
1
2
3

## Slide 24
1
2
3
+5V
0V
-5V
+5V
0V
-5V
+5V
0V
-5V
1
2
3

## Slide 25
1
2
3
+5V
0V
-5V
+5V
0V
-5V
+5V
0V
-5V
1
2
3

## Slide 26: CCD Performance Metrics
Charge generation
quantum efficiency (QE), dark current
Charge collection
full well capacity, pixels size, pixel uniformity, defects, diffusion (Modulation Transfer Function, MTF)
Charge transfer
charge transfer efficiency (CTE), defects
Charge detection
readout noise (RON), linearity

## Slide 27: Well Capacity
Blooming
pixel
boundary
Photons
Photons
Overflowing
charge packet
Spillage
Spillage
pixel
boundary

## Slide 28: Saturation
CCD pixels have a linear response of measured output voltage to a value quite close to the full well capacity of the pixel. The number of electrons which can be stored is given by:
Q = CV
V is the voltage, and C is the capacitance of the pixel, given approximately by:
C  Aκε0/d
A is the area of the pixel, d is the thickness of the SiO2 layer, κ is the dielectric constant of SiO2 (about 4.5) and ε0 is the permittivity of free space.
Notes: 28
We look at the charge capacity or full well capacity of a CCD pixel. CCds have a linear response, I.e. charge is proportional to radiation, to close to this capacity. We can work out the capacitance of a pixel and thus thus the full well capacity. A few things to note, the capacity of proportional to the area of the pixel, so our drive to get smaller pixels for better resolution has a drawback. The capacity is also proportional to the Voltage, MPP chips have a lower voltage difference between the central phase and surrounding phases.

## Slide 29: Linearity and Saturation
Typically, the full well capacity of a CCD pixel 25 μm square is 500,000 electrons.
  If the charge in the well exceeds about 80% of this value the response will be non-linear.
  This is because the filled pn junction region will have a weaker and weaker field.
  If it exceeds this value, charge will spread through the barrier phase to surrounding pixels.
This charge blooming occurs mainly vertically, as there is little horizontal bleeding because of the permanent doped channel stops.
Readout register pixels are larger, so there is less saturation effect in the readout register.
Notes: 29
We can work this number out, and there is a problems on this. If we exceed the capacity then charge tends to spread up and down the columns. Not across because of the permanent channel stops, and not in the readout register where we increase the pixel area to increase the charge capacity.

## Slide 30: Blooming
Blooming
Bloomed star images

## Slide 31: Dark Current
Dark current is generated when thermal photon-induced vibrations cause an electron to move from the valence band to the conduction band.
The majority of dark current is created near the interface between the Si and the SiO2, where interface states at energy between the valence and conduction bands act as a stepping stone for electrons. These are sometimes referred to as “mid-band” states.
CCDs can be operated at temperatures down to around 140K, to reduce thermal effects.
Notes: 31
The dark current is background signal generated by thermal effects. Because of the dark current CCds are run cooled, to reduce the possibility of thermal excitation of electrocs across the band gap.

## Slide 32: Dark Current vs. Temperature
Thermally generated electrons are indistinguishable from photo-generated electrons : “Dark Current” (noise)
Cool the CCD down!!!

## Slide 33: Defects: Dark Columns
Dark columns: caused by ‘traps’ that block the vertical transfer of charge during image readout.
Traps can be caused by crystal boundaries in the silicon of the CCD or by manufacturing defects.
Although they spoil the chip cosmetically, dark columns are not a big problem (removed by calibration).

## Slide 34: Defects: Bright Columns
Cosmic rays
Cluster of
hot pixels
Bright
column
Bright columns are also caused by traps. Electrons contained in such traps can leak out during readout causing a vertical streak.
Hot pixels are pixels with higher than normal dark current. Their brightness increases linearly with exposure times
Somewhat rarer are light-emitting defects which are hot pixels that act as tiny LEDS and cause a halo of light on the chip.

## Slide 35: Charge Transfer Efficiency
When the wells are nearly empty, charge can be trapped by impurities in the silicon. So faint images can have tails in the vertical direction.
Modern CCDs can have a charge transfer efficiency (CTE) per transfer of 0.9999995, so after 2000 transfers only 0.1% of the charge is lost.
Notes: 35
Charge can be trapped as you read it out, although with modern buried channel CCDs this is much less of a problems than it used to be.

## Slide 36: Charge Transfer Efficiency
CTE = Charge Transfer Efficiency (typically 0.9999 to 0.999999)
= fraction of electrons transferred from one pixel to the next
CTI = Charge Transfer Inefficiency = 1 – CTE (typically 10– 6 to 10– 4)
= fraction of electrons deferred by one pixel or more
Cause of CTI:
charges are trapped (and later released) by defects in the silicon crystal lattice
CTE of 0.99999 used to be thought of as pretty good but ….
Think of a 9K x 9K CCD
Notes: 36

## Slide 37
Example:
X-ray events with charge smearing in an
irradiated CCD (from GAIA-LU-TN01)
direction of charge transfer
In the simplest picture (“linear CTI”) part of the
original image is smeared with an exponential
decay function, producing “tails”:
original image
after n transfers
37
Notes: 37

## Slide 38: CTE
Percentage of charge which is really transferred.
“n” 9s: five 9s = 99,99999%

## Slide 39: Read-Out Noise
Mainly caused by thermally induced motions of electrons in the output amplifier.
These cause small noise voltages to appear on the output.
This noise source, known as Johnson Noise, can be reduced by cooling the output amplifier or by decreasing its electronic bandwidth. Decreasing the bandwidth means that we must take longer to measure the charge in each pixel, so there is always a trade-off between low noise performance and speed of readout.
The graph below shows the trade-off between noise and readout speed for an EEV4280 CCD.

## Slide 40: CCD readout noise
Reset noise: there is a noise associated with recharging the output storage capacitor, given by σres=  (kTC) where C is the output capacitance in Farads.
This is removed by correlated double sampling, where the reset voltage is measured after reset and again after readout. The first value is subtracted from the second, as this voltage will not change.
The output Field Effect Transistor also contributes noise. This is the ultimate limit to the readout noise, at a level of 2-3 electrons
Notes: 40
Surface state noise is akin to charge transfer problems, and like those is much reduced in buried channel CCDs.
The largest source of noise in the best modern CCDs is the FET noise from the previous slide.

## Slide 41: kTC Noise
Consider the thermal equipartition of charge on a capacitor.
Boltzmann statistics apply
We are counting charges, so the variance is given by Poisson statistics.

## Slide 42: Other noise sources
Fixed pattern noise. The sensitivity of pixels is not the same, for reasons such as differences in thickness, area of electrodes, doping. However these differences do not change, and can be calibrated out by dividing by a flat field, which is an exposure of a uniform light source.
Bias noise. The bias voltage applied to the substrate causes an offset in the signal, which can vary from pixel to pixel. This can be removed by subtracting the average of a number of bias frames, which are readouts of zero exposure frames. Modern CCDs rarely display any fixed pattern bias noise
Notes: 42
Fixed pattern noise can be quite serious, but is calibrated out with a flat field.

## Slide 43: Interference Fringes
In thinned CCDs there are interference effects caused by multiple reflections within the silicon layer, or within the resin which holds the CCD to a glass plate to flatten it.
These effects are classical thin film interference (Newton’s rings).
Only visible if there is strong line radiation in the passband, either in the object or in the sky background.
Visible in the sky at wavelengths > 700nm.
Corrected by subtracting off a scaled exposure of blank sky.
Notes: 43
Fringing can dominate the noise in the redder photometric bands, or in narrow bands, and can sometimes force us back to using thick CCDs despite the loss in RQE.

## Slide 44: Examples of fringing
Fringing on H1RG SiPIN device at 980nm
Notes: 44

## Slide 45: Correcting Fringing from Sky Lines
Notes: 45
CCDs are made of silicon, which becomes transparent with increasing wavelength, starting at around 700nm. In back-illuminated, thinned CCDs airglow can cause internal reflections leading to a characteristic interference pattern in the background of an image, known as fringing. The pattern itself depends mainly on the thickness of the CCD and is thus essentially unchanged over the lifetime of the detector. However, the amplitude of the fringing pattern varies as the airglow spectrum at a given position on the sky changes with time. Fringing is also present in near-IR detectors.

## Slide 46: Large CCD Mosaics

## Slide 47: LSST Has a Big Camera

## Slide 48: LSST Has a Big Focal Plane
Guide Sensors (8 locations)
Wavefront Sensors (4 locations)
3.5 degree Field of View (634 mm diameter)

## Slide 49: LSST CCD Raft

## Slide 50: Basic CCD Camera
.
.
.
Thermally Electrical feed-through Vacuum Space Pressure vessel Pump Port
Insulating
Pillars
Focal Plane
of Telescope
Telescope beam
Optical window CCD CCD Mounting Block Thermal coupling Nitrogen can Activated charcoal ‘Getter’
Boil-off
Face-plate

## Slide 51: CCD calibration
Flat Field Image
Bias Image
Flat
-Dark
-Bias
Science
-Dark
-Bias
Output Image
Flat-Dark-Bias
Sc-Dark-Bias
Dark Frame
Science Frame
If there is significant dark current present:

## Slide 52: CCDs for X-ray Applications
Notes: 52
Fringing can dominate the noise in the redder photometric bands, or in narrow bands, and can sometimes force us back to using thick CCDs despite the loss in RQE.

## Slide 53: infrared detectors

## Slide 54: History
Herschel’s detection of IR from Sun in 1800
Johnson’s IR photometry of stars (PbS) mid 60’s
Neugebauer & Leighton: 2um Sky Survey (PbS), late 60’s
Development of bolometer (Low) late 60’s
Development of InSb (mainly military) early 70’s
IRAS 1983
Arrays (InSb, HgCdTe, Si:As IBCs) mid-80’s
NICMOS, 2MASS, IRTF, UKIRT, KAO, common-user instruments, Gemini, etc.
JWST and the search for cosmic origins

## Slide 55: Historical motivation
Exploration & discovery
  Neugebauer, Leighton, Low, Fazio, Townes
Technological opportunities
  Bolometer (Low)
  PbS (Neugebauer)
  balloons (Fazio)
  IR lasers & interferometry (Townes)
A few, key problems
  Bolometric luminosities (Herschel, Johnson)
  The Galactic Center (Becklin)
  Star formation

## Slide 56: Current interest in infrared
High redshift objects
  lobs = l0 (1+z) 5000 Å ® >1 mm for z > 1
  Classical problems require infrared data
Obscuration by dust
  Al ~ l-1.9 ® A2.2mm ~ 0.1 AV (Mathis 1990, ARAA, 28, 37)
  Now important for:
    Galactic nuclei, esp. AGN (unified model)
    Starburst galaxies
    Young stars
Very low mass objects & extrasolar planets
  Tplanet ~ 50 to 500 K lpeak ~ 5 – 50 mm
  TBD ~ 900 – 2000 K lpeak ~ 1 – 5 mm

## Slide 57: Extinction by dust
  J. S. Mathis 1990, ARAA, 28, 37.

## Slide 58: Infrared Hybrid Array
Infrared arrays use light-sensitive material that can detect infrared photons, wavelengths beyond ~1 um.
Therefore, silicon cannot be used as the light-sensitive layer.
This poses a problem because the readout circuit is most easily implemented in silicon.
Therefore, infrared arrays are “hybrids” – they use one material to detect light and silicon for the readout circuit.

## Slide 59: Indium Bumps

## Slide 60: Pixel-level Cross Section: InSb

## Slide 61: Pixel-level Cross Section: HgCdTe

## Slide 62: NICMOS

## Slide 63: Si PIN Architecture
Si PIN devices are architecturally similar to infrared arrays, so we discuss them in this lecture even though they do not detector infrared light (they use silicon!).
A PIN is a sandwich of p-type/instrinsic/n-type layers.
These devices are usually implemented for improved response at longer wavelengths, i.e. they are thick.
In order to ensure that photogenerated charge makes it to the integrating node, there needs to be a strong electric field that depletes the intrinsic layer.
PIN diode

## Slide 64: Detector Size

## Slide 65: Valence & Conduction Bands in Semiconductors
When atoms (a) come together to form a crystal, the outer energy levels overlap and blend to create bands (b).
The outermost filled band is called the valence band (c).
Above the valence band, one finds a forbidden energy gap -the “band gap”, and (at higher energies) conduction bands populated by thermally excited electrons.
In metals, the valence and conduction bands overlap resulting in conduction. In insulators, the band gap is wider resulting in very poor conduction.

## Slide 66: Periodic Table
Semiconductors occupy column IV of the Periodic Table
Outer shells have four empty valence states
An outer shell electron can leave the shell if it absorbs enough energy

## Slide 67: Periodic Table Continued
The column number gives the number of valence electrons per atom. Primary semiconductors have 4.
Compounds including elements from neighboring columns can be formed. These alloys have semiconductor properties as well (e.g. HgCdTe & InSb).
Mercury-cadmium-telluride (HgCdTe; used in NICMOS) and indium-antimonide (InSb; used in SIRTF) are the dominant detector technologies in the near-IR.

## Slide 68: The Band Gap Determines the Red Limit

## Slide 69: NICMOS MUX

## Slide 70: Pixel Capacitance and Gain

## Slide 71: Dark Current
Like a CCD, IR arrays are affected by dark current and a variety of noise mechanisms.
Dark current is the signal that is seen in the absence of any light. For the near-IR (1-2.5 μm), the dominant components are diffusion, thermal generation-recombination (G-R) of charges within the semiconductor, and leakage currents. Combining these, it can be shown that where V is the voltage across the detector and R0diff and R0GR are the detector impendences at zero bias for diffusion and G-R.

## Slide 72: Dark Current Example

## Slide 73: Dark Current vs. Temperature

## Slide 74: kTC Noise
“kTC” noise occurs in both CCDs and IR arrays when the detector capacitor is recharged. As we will see shortly, this goes as
where k is the Boltzmann constant, T is temperature, C is capacitance, and e is the elementary charge.

## Slide 75: 1/f Noise
1/f noise in the multiplexer’s output field effect transistor (FET).
  In the mid-1980s, readout noise ~300 electrons was common. Better output FETs increased the number of microvolts per electron and thereby helped to reduce readout noise to <10 electrons seen today.
Other 1/f components are seen. e.g.,
  The “pedestal effect” seen in NICMOS whereby the entire exposure’s bias is seen to vary by ~75 electrons. There is some consensus that these 1/f components are thermally driven.

## Slide 76: Noise Reduction Through Multiple Sampling
By being a bit clever about reading out the array, one can minimize or eliminate some of these noise modes.
During an exposure, typically each pixel is sampled several times.
The most common approaches are correlated double sampling (CDS), multiple non-destructive reads (aka “Fowler Sampling”), & fitting a line (aka “up the ramp”). We discuss these briefly.

## Slide 77: Fowler Sampling

## Slide 78: “Up the Ramp”
Fit best line to multiple non-destructive samples.
Sample spacing does not need to be uniform.
Not clear whether this or Fowler sampling is best.
This is what is done in NICMOS MULTIACCUM mode.

## Slide 79: Read Noise Example

## Slide 80: QE Example
Relative QE for Bepi-Colombo device, at 293 K, of RVS MCT array after substrate removal. High response in the UV region is due to charge gain. The QE reduces by ~8% at 190 K.

## Slide 81: Well Depth and Non-linearity

## Slide 82: Persistence
Persistence is the tendency of a detector to capture charge liberated in one exposure and then release it later in another exposure.
It is produced by charge traps.
Persistence introduces a noise source into images.

## Slide 83: Persistence Continued

## Slide 84: Persistence Movie

## Slide 85: Pixel-to-pixel Crosstalk
Crosstalk occurs when a pixel value is influenced by its neighbors.
Charge diffusion is an important crosstalk mechanism in IR arrays and CCDs.
Once charge carriers are created, their motion is governed by charge diffusion.

## Slide 86: IPC
Interpixel capacitance (IPC) is another form of crosstalk.
In this case, charge in a pixel induces a voltage change in a neighbor, just like the behavior between parallel plates in a capacitor.
The effect is to blur the point spread function.
The induced voltage does not have noise.

## Slide 87: IPC
In this example, IPC is very large for the H4RG SiPIN device (10 um pixel size).

## Slide 88: Teledyne Family Arrays
H2RG HgCdTe array

## Slide 89: Teledyne H4RG Si PIN
This is an image of the H4RG device in the Rochester Imagingn Detector Laboratory.

## Slide 90: Raytheon Family Arrays

## Slide 91: MCT on Si
Design of the light-sensitive part of the pixel for the RVS SB-390 MCT on Silicon (Bornfreund et al. 2005).

## Slide 92: MCT on Si