## Slide 1: Development of High Z Photon-Counting X-ray Detectors for Synchrotron Experiments Above 30 keV
M. Ruat, C. Ponchut
Scientific Detector Workshop, Firenze, October 10th 2013

## Slide 2: Outline
Context
High Z detectors developments at ESRF
Bottleneck(s) and improvements
Beamline evaluations
Future plans

## Slide 3: Outline
Context
High Z detectors developments at ESRF
Bottleneck(s) and improvements
Beamline evaluations
Future plans

## Slide 4
Detector developments for synchrotrons
Synchrotron == BRILLIANCE!
WIDE RANGE OF APPLICATIONS
Macromolecular crystallography (MX)
Single crystal diffraction (SCD)
Surface diffraction
Small- and wide-angle X-ray scattering (SAXS/WAXS)
Coherent X-ray imaging
Pre-clinical X-ray imaging
Time-resolved experiments
1022 B
Up to 1013 photons / mm2/s

## Slide 5
2D Detectors
In house
In house
In house
Commercial
Commercial
Commercial

## Slide 6
Direct Detection Detectors
CCD/CMOS

## Slide 7
Direct Detection Detectors
CCD/CMOS

## Slide 8
Single photon counting with threshold: noise-free detection
Direct Detection Detectors  Single Photon Counting
Today all direct detection & single photon counting detectors are equipped with Si sensors

## Slide 9
Towards High Z sensors
ESRF UPGRADE:
Increased brilliance & increased photon energies !!

## Slide 10
    2D detector
    ~100 ns time resolution
    kHz frame rates
    High Z sensor layer
Direct detection: Filling gaps
    2D APD array detector
    sub-ns time resolution
    MHz frame rates
  Photon counting
  pixel detectors
    Mpixel devices
    kHz frame rates
    ~100 ns time resolution
  MAXIPIX
  1500 frames/sec
  APD
  avalanche photodiodes
    Point or linear devices
    Sub-ns time resolution
  ACE
  5 ns deadtime
  ~100 MHz count rate
  Towards harder X-rays
  Towards higher time resolution
1D detectors   2D detectors
Towards High Z sensors
Notes: 10

## Slide 11: Outline
Context
High Z detectors developments at ESRF
Bottleneck(s) and improvements
Beamline evaluations
Future plans

## Slide 12
HIZPAD project
Development in the framework of HIZPAD European project (started in 2009)
Coordinated by ESRF

## Slide 13
Why CdTe?
CdTe 1 mm
Si 500 µm
GaAs 300 µm
+ Wide, direct band-gap (1.44 eV) Room temperature operation
Notes: 13

## Slide 14
Hybrid Pixel Detector
CMOS RO ASIC
Sensor: CdTe 1 mm thick
ASIC: from MEDIPIX family (CERN)  55 um pixel pitch
Interconnects: Bump-bonding (In-Sn)

## Slide 15: HIZPAD at ESRF  Maxipix-CdTe, a direct detection imaging system for synchrotron beamlines in the 5 to 150 keV energy range
Applications:
Material science:
Surface diffraction, High-resolution diffraction, time-resolved diffraction, diffraction tomography, ptychography, …
X-ray Imaging
Sensor material | CdTe, single crystal (Acrorad)
Pixel pitch | 55 μm
Sensor size (pixels) | 512 x 512 (ohmic type)
Sensor size (mm) | 28 x 28 x 1
X-ray detection mode | Single photon counting
ASIC | TIMEPIX (CERN)
Readout | ESRF MAXIPIX, max speed 1.4 kHz
Flux (X/mm2/s) | Up to 3.107 (Timepix limited)
Operation Temperature | Room Temperature
Absorption efficiency | 0.99 up to 70 keV
Current detector system : Presentation

## Slide 16
M. Ruat and C. Ponchut, IEEE Trans. Nucl. Sci., Vol. 59, n° 5, pp. 2392-2401, August 2012
C. Ponchut and M. Ruat, Proceedings of IWORID 2012
M. Ruat, E. Gros d’Aillon and C. Ponchut, proceedings of the 2012 IEEE NSS/MIC/RTSD conference
Spatial resolution 63 um
Linear count rate up to 3.107 photons/mm2/s
Yb2O3 nano-powder diffraction rings at 50 keV.
Comparison with FRELON (50 um pixel pitch) :No significant background, sharper peaks, more low intensity rings
Current detector system : Performances

## Slide 17: Outline
Context
High Z detectors developments
Bottleneck(s) and improvements
Beamline evaluations
Future plans

## Slide 18
Small pixels enables selection of fast electrons (single carrier sensing) but increase of charge sharing phenomenon
Charge sharing and fluorescence
Mean free path of fluorescence photons : 110 um
 They escape from the illuminate pixel and contribute to charge sharing
Only 50 % of the pixel area is not shared with neighbors  double counting
Escape peak
 Corrections at the ASIC level (Charge summing ASIC)
Fluorescence
PSF

## Slide 19
1mm
1mm
Flat-field image of a CdTe detector:
Observation of structural defects

## Slide 20
1mm
Reduced color scale around median value
Flat-field image of a CdTe detector:
Observation of structural defects

## Slide 21
*
*
1mm
1. Network of Lines (Native):
Increased counts by ~10% next to a line of reduced counts (drift of the charges to the neighbor pixels)
Dislocations walls / grain boundaries
Can be FF corrected
M. Ruat and C. Ponchut, IEEE Trans. Nucl. Sci., Vol. 59, n° 5, pp. 2392-2401, August 2012
C. Ponchut and M. Ruat, Proceedings of IWORID 2012
M. Ruat, E. Gros d’Aillon and C. Ponchut, Proceedings of the 2012 IEEE NSS/MIC/RTSD conference
bump-bonding yield of 98 to 99%
*
*
* From C. Buis et al, Proceedings of the 2012 IEEE NSS/MIC/RTSD conference) 2011 IEEE, pp. 4804 - 4808
Flat-field image of a CdTe detector:
Observation of structural defects

## Slide 22
1mm
2. Extended clusters of saturated pixels (from bump-bonding process):
- current overflow of the analog input stage of the readout chip.
- Strain in the material (From E. Hamman et al, IEEE RTSD 2012)
Surrounded by large areas of reduced counts (gettering effect)
On-going development for improvement
M. Ruat and C. Ponchut, IEEE Trans. Nucl. Sci., Vol. 59, n° 5, pp. 2392-2401, August 2012
C. Ponchut and M. Ruat, Proceedings of IWORID 2012
M. Ruat, E. Gros d’Aillon and C. Ponchut, Proceedings of the 2012 IEEE NSS/MIC/RTSD conference
Flat-field image of a CdTe detector:
Observation of structural defects

## Slide 23: Some of the main milestones of the CdTe community:
- Quality of the material with respect to structural defects
- Contacting and bonding
3D rendering of CdTe on Timepix failed bump-bonding (pixel pitch 110 um)
From laminography experiment, ESRF - ID15, June 2012
Acknowledgments to A. Cecilia and E. Hamman from ANKA
Bump-bonding

## Slide 24
Correction for structural defects / bump-bonding faults
However….. These are active defects  Temporary correction!!

## Slide 25
1. Irradiation with X-ray tube (Ag anode, Kα 22 keV)
7.105 photons/mm2/s for 7 hours (1.8e10 ph./mm2)
With lead pattern in front of the detector
Last image taken with lead pattern removed
 Phantom image of the lead pattern due to trapped charges
M. Ruat, proceedings of the IWORID 2013
2. Irradiation at BM05, E = 15 keV
Left part of the irradiated zone: 105 photons/mm2/s
Right part 1010 photons/mm2/s, for 10h (3.6e15 ph./mm2)
 Lateral polarization == local electric field distortions due to differential irradiation
Instabilities : charging of the sensor
3 mm thick Lead plate

## Slide 26: Irradiation with 30 keV X-rays @ 1010 photons/mm2/s for 12h
Before Irradiation
After 6h of continuous irradiation (right part)
After bias reset
Polarization / charging can be avoided (to a certain extent)!
 Need to improve the design:
- Cooling or controlling the operation temperature and / or
- Reset the bias applied to the detector at regular interval times (known technique for CdTe detectors with Schottky diodes contacts)
Instabilities : charging of the sensor

## Slide 27
Dark and Flat-field images: as is when received
After 12h at 80°C (new noisy pixels on dark image existed before baking)

## Slide 28: Outline
Context
High Z detectors developments
Bottleneck(s) and improvements
Beamline evaluations
Future plans

## Slide 29: Beamline evaluations at ESRF
ID15 – Surface diffraction: Reflectivity liquid Ga on solid Sapphire at 70 keV (Oct. 2012)
ID15 – Surface diffraction: In-situ following of GaN growth at 70 keV (Oct. 2012)
ID11 – High resolution diffraction imaging at 80keV. Mechanical stability of retained austenite grains in TRIP steels studied by synchrotron X-ray diffraction during deformation (Nov. 2011)
ID17 - Edge-Illumination phase contrast imaging for mammography at 60 and 85 keV (Sept. 2012)
ID22 - Coherent SAXS or Ptychography experiment at 20 keV (Sept. 2012)
ID15 – Fast SAXS-WAXS experiment with Perkin Elmer detector for WAXS and Maxipix-CdTe as SAXS detector
ID17 – Detector evaluation for new tomography setup
for paleontology (June 2013)
ID17 – Dosimetry reconstruction for MRT (July 2013)

## Slide 30: Outline
Context
High Z detectors developments
Bottleneck(s) and improvements
Beamline evaluations
Future plans

## Slide 31: FUTURE PLANS
Detector design improvement to handle for CdTe material specific requirements
 Reset bias routine, Controlled operation temperature, improve contacts and bonding (HIZPAD2)
Charge sharing effect compensation
 Medipix 3 ASIC
Large area and/or modular architecture, requires edgeless sensor and 4-sides buttable ASIC
 HIZPAD2
Faster counting and faster frame rate
 Medipix 3, Eiger, Pilatus 3
Alternative sensor materials e.g. GaAs

## Slide 32
HIZPAD2 is part of a JRA starting 2013 for 3 years
Budget: 1 Million euros
Partners: ELETTRA, CNRS/CPPM, CNRS/Néel, DESY, DLS, PSI, SOLEIL, STFC and EXFEL as observer.
Coordinated by ESRF
Dedicated to:
  Improved sensor quality and processes (specifically contacts & bump-bonding)
  Edgeless detectors and bigger pixel arrays
  Build enhanced demonstrators for beamline evaluations
  Theoretical approaches (modelling)
  Explore innovative X-ray sensor concepts, i.e. Quantum well materials

## Slide 33: Many thanks to the following people who contributed to this presentation: The ESRF Detector Unit Collaborators of the ESRF Science Division Our partners from European projects Thank you for your attention!

## Slide 34: ESRF Development "roadmap"
2005
2010
2015
MAXIPIX (Si sensor)
1280x256 pix./300 fps, 256x256 pix./1400 fps
SMARTPIX
>2000 fps, energy resolution, 1k x 1k
high-Z pixel sensors (HIZPAD, HIZPAD2)
high energy applications
nb. applications
14 units in operation at ESRF
MAXIPIX for UPBL
specific implementations
Pixel sensors improvements
active edge for larger areas
EIGER (collab. PSI)
>20 000 fps, large area
Context: Photon counting hybrid Pixel Detectors
faster
Energies> 30 keV
Large area
Energies< 20-30 keV

## Slide 35: Irradiation with 30 keV X-rays @ 1010 photons/mm2/s for 12h
Before Irradiation
After 6h of continuous irradiation (right part)
Immediately after bias reset
Bias reset + 10 min
Bias reset + 30 min
Note: Left part of the illuminated zone was irradiated at flux within the linear counting range: no measurable charging effect

## Slide 36: Beam viewer
Current developments: Hybrid Pixel Detectors
M. Ruat - 3-WAY Meeting Aug. 2013 - APS
TECHNOLOGY | CURRENT | FUTURE
SOFT X-RAYS LARGE FIELD OF VIEW NEEDED | Si sensors TIMEPIX chip | Edgeless Si sensors - Medipix 3 & SMARTPIX
HARD X-rays | CdTe sensor TIMEPIX chip | - HIZPAD2 project: Slim-edge High-Z sensors, Improved processing - Medipix3 chip
The MAXIPIX Detector and its evolution

## Slide 37
Longer beamlines, increased photon energy, increased flux, new techniques  New detector solutions
Smaller objects  less signal  noise free detectors+ spatial resolution
Synchrotron is providing high spatial resolution  you want to have it on the detector as well (optics solution + direct detection solutions)
Context: ESRF Upgrade

## Slide 38
Phantom images of previous irradiation field
Instabilities : temporary radiation damage

## Slide 39
Soleil synchrotron 20 keV + air scatterer
Thld 5
12 x 60 s
Flat-Field when received (2012-02)

## Slide 40
Ag Tube 35kVp – 25 mA
+ 100 um Ag filter
Thld 11
100 x 1 s
As is, after 18 months of experiments…
Chips 00, 01, 11 exhibit more extended defects + history effect.
Chip 10 is noisy and some pixels look disconnected

## Slide 41
After 2 hours at 50°C…
Ag Tube 35kVp – 25 mA
+ 100 um Ag filter
Thld 11
60 x 1 s
History effects have disappeared, as well as some defects on chips 00, 01, 11. Chip 10 idem  temporary :-(

## Slide 42
Irradiation with or without bias reset every 10 minutes
Irradiation with X-ray tube (Ag anode, Kα 22 keV, 35 kVp, 50 mA)
~1.107 photons/mm2/s for 7 hours on right part of the detector (7.2e10 ph./mm2)
Left part of the irradiated zone < 103 photons/mm2/s
 Imaging a lead pattern before and after irradiation
103 ph./mm2/s
107 ph./mm2/s
2 mm thick Lead plate
Reset bias routine

## Slide 43
Irradiation with or without bias reset every 10 minutes
Irradiation with X-ray tube (Ag anode, Kα 22 keV, 35 kVp, 50 mA)
~1.107 photons/mm2/s for 7 hours on right part of the detector (7.2e10 ph./mm2)
Left part of the irradiated zone < 103 photons/mm2/s
 Imaging a lead pattern before and after irradiation
Without bias reset
Reset bias routine

## Slide 44
Irradiation with or without bias reset every 10 minutes
Irradiation with X-ray tube (Ag anode, Kα 22 keV, 35 kVp, 50 mA)
~1.107 photons/mm2/s for 7 hours on right part of the detector (7.2e10 ph./mm2)
Left part of the irradiated zone < 103 photons/mm2/s
 Imaging a lead pattern before and after irradiation
Without bias reset
With bias reset
Reset bias routine

## Slide 45
Polarization / charging
Extensive use  quicker polarization, enhanced radiation damage, extra noise
Low temperature baking (50°C) induces defects relaxation in CdTe
NEED TO CHARACTERIZE TEMPORARY vs PERMANENT or RECURRING EFFECTS
And the way to avoid or ‘anneal them’, if possible.
 Need to improve the design:
- Cooling or controlling the operation temperature and / or
- Reset the bias applied to the detector at regular interval times (known technique for CdTe detectors with Schottky diodes contacts)
Detector instabilities: Overview

## Slide 46: ID15 – Surface diffraction on liquid Ga(N) / solid Sapphire interface at 70 keV
In-house experiment with A. De Jong, J. Jasper and V. Honkimaki from ID15; Pr. E. Vlieg from Radboud University Nijmegen
Main outcomes from the use of 2D CdTe detector:
Faster and simplified sample and motors alignment procedure
High energy needed for buried interfaces studies
Measurement speed improvement:
Study fluctuations of intensity between Bragg peaks need for a sensitive detector with high dynamic range
72h of continuous irradiation without significant drift

## Slide 47: ID17 – Edge-Illumination phase contrast imaging for mammography at 60 and 85 keV
In-house experiment with A. Olivo and P. Diemoz from University College London; P. Coan from Ludwig Maximilians University Munich; A. Bravin and T. Brochard from ID17
From P. Munro et al., AIP Conf. Proc. 1466, 118 (2012)
Method
Results
Simplified alignment
Ability to work at high energies
Spatial resolution ok
Identification of small calcified zones due to tumor into human breast (60 keV)
8m
slits
Illuminated zone 512 x 1 pixel
48h of continuous irradiation without significant drift
2 bad lines due to defective pixels have been removed

## Slide 48: ID11 - Mechanical stability of retained austenite grains in TRIP steels studied during deformation
Users: Romain Blondé et al., Delft University of Technology, The Netherlands; in collaboration with J. Wright ID11
From R. Blondé et al., AIP Conf. Proc. 1466, 118 (2012)
Method : High resolution X-ray diffraction
TRIP  TRansformation Induced Plasticity effect.
Monitoring of individual grains in the bulk of the material to follow the formation of sub-grains prior to the transformation during deformation.
Conclusion: High-energy micro-beam XRD is a powerful tool to explore advanced multiphase steels

## Slide 49: ID22 – Coherent SAXS or Ptychography experiment at 20 keV First commissioned CdTe detector (single chip)
High number of photons in the center, CdTe needed to protect the readout chip
Low number of photons outside the center: photon-counting feature needed
Additional detector features needed: bad pixel masking and dead time correction to be accurately quantitative
Setup 1 (FRELON like):
1. Low flux scan
2. High flux scan
Setup 2 (50% gain in time):
Only one scan, attenuator in the center
1
2
2
First experiment with Maxipix-CdTe done in Sept. 2012 with P. Thibault, F. Pfeiffer et al. from Technische Universität München, Germany and P. Cloetens ID22:
Object: Siemens Star
Object: Nano-porous Au

## Slide 50: ID22 – Coherent SAXS or Ptychography experiment at 20 keV First commissioned CdTe detector (single chip)
Integrating: 'Infinite' count rate
Counting: High sensitivity CdTe Maxipix
Future UPBL4 detector
 Future developments: UPBL4

## Slide 51
Flat-fields:
A few additional disconnected pixels, not in the ‘damaged’ zone
Some blobs decreased in size (o), some increased in size (o)
Some blobs have some ‘rays’ linked to new structures (o)
Scratch or other undefined features linked to new structures? (--, o)

## Slide 52
Lead pattern image: ‘Vamos a la playa’
Comes from the backside contact?
Comes from the CdTe bulk crystal?
Comes from the pixel contacts?

## Slide 53
Stretched and Rotated image
Flat-Field
TOPOGRAPHY
Bragg geometry, grazing incidence, 20 keV monochromatic beam, 004 reflection
Can see blobs and dislocation lines. Faint lines looking like additional structures. Blobs aligned with these lines?
Large vertical structures visible with natural light on backside contact

## Slide 54
EBM, 5 keV
EBM, 15 keV
BACKSIDE CONTACT SURFACE STATE
AFM
DIC microscopy
Conventional microscopy
EBM