## Slide 1: In situ testing of detector controllers
Roger Smith
Caltech
2013-10-10

## Slide 2: The need
Many “detector problems” are actually controller failures.
High channel counts make testing expensive:
  32 channels are common for NIR detectors
  Many CCD cameras are high channel count mosaics.
How to test with minimal human intervention?
This talk is an attempt to share some useful ideas rather than to be a comprehensive diagnostics manual.
Will focus on automated noise and DNL tests

## Slide 3: Example problems
ADC DNL (e.g. missing codes)
Excess readout noise or bias drift (banding)
  Component failure (rare; usually bypass capacitors)
  Noisy power supplies
  Noisy bias voltages.
  ADC or DAC reference voltage drift.
  Bad grounding/shielding
Clock failures (wiring, connectors?)
Supply filtering failure; especially bypass caps.
Conducted interference (from AC outlet) or in shields (from telescope)
Non-Linearity, usually clipping, signal offsets.

## Slide 4: SENSORY OVERLOAD ???
…..Let’s focus on noise and ADC DNL

## Slide 5: Simplified block diagram
Detector
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 6: What test equipment ?
To see the noise and signals of interest we need
Low input referred noise and differential probe.
Same passband as the CDS processor
Samples to be synchronized to pixels.
  (Interference that is same on every pixel is benign)
 Controller is its own oscilloscope!
Typical scope input ~ few mV (20Mhz BW limit)
CCD output noise ~ 3e- * 5µV/e- = 15µV

## Slide 7: Data link and pixel sorting
Clk DAC
Clk DAC
Bias DAC
ADC
Artificial Data
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 8: Will discuss ADC testing later…
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 9: Permanent integrator reset
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 10: Permanent (black level) clamp
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 11: Shorted input – loopback connector
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 12: Bias loopback at connector or mount
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield
Short

## Slide 13: Bias DAC test
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 14: Clock loopback – stationary low
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 15: Clock loopback – stationary high
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 16: Clock loopback – stationary high
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 17: Power Supply loopback
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 18: Shield loopback -- make connection near detector if possible
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 19: CCD Reset feedthru  switching threshold
Detector
Clk DAC
Clk DAC
Bias DAC
ADC
Data Link
Host
Clock
Bias
Preamp
Noise BW control
Optical Fiber
Disk
Power supplies
AC coupler (for CCDs)
Network
RFI shield

## Slide 20: ADC DNL

## Slide 21: What is Differential Non-Linearity?
Common causes:
Differences between analog and digital ground at ADC.
Pixel synchronous transient on ADC supply or reference.
Inductance in bypass cap leads; self inductance.
Reference or signal source impedance too high at bit rate.
Variation in voltage delta corresponding to one output code, usually expressed in fraction of nominal.
+1ADU implies that code represents twice the usual input range
Poor DNL can suppress or exaggerate noise depending on signal level.

## Slide 22: How to measure DNL?
Too hard to scan input voltage to look for thresholds directly (noise, drift), so instead…
Generate input signal with smooth histogram
  flat is good but not required.
Make histogram of raw data. Must be totally unprocessed.
Counts per bin are proportional to voltage delta represented by each code.
  Works just as well in the presence of noise !
  Need enough data for statistical errors to be negligible compared to fixed pattern.
    Eg: 1000samples_per_bin * 216 bins * 3µs/sample = 196 sec.

## Slide 23: Measure all channels at once

## Slide 24: Histograms of DNL errors
All these are acceptable, but performance varies greatly…..

## Slide 25: Zooming in on poor ADC

## Slide 26: Excellent ADC
Fourier Transform of DNL errors reveals that they occur at binary transitions as expected.
Histogram of DNL errors

## Slide 27: Rasterize histogram to see details

## Slide 28: Rasterize histogram to see details

## Slide 29: Rasterize histogram to see details

## Slide 30: Rasterize histogram to see details

## Slide 31: Making flat input histogram with CCD
Erase with shutter closed then read CCD with shutter open to generate linear signal ramp
  Read noise, shot noise and flat field irregularities are ok.
Adjust signal and/or line time so full well reached at end end of readout.
Combine histograms from multiple images. Need ~65 megapixels per output amplifier so ~ 1000 counts DAC code (assuming 16 bit ADC).

## Slide 32: Generating flat input histogram with CMOS mux
Use unsubtracted data (not CDS)…
Pixel offsets are normally distributed with typically ~3000 ADU FWHM.
Mean can be shifted by input offset DAC.

## Slide 33: Overlapped gaussians  flat top
In our NIR detector controllers (ARC Inc), the input offset DAC produces shift in mean ~130 ADU at ADC for 1 count at DAC. This is safely less than 1500 ADU FWHM
Gaussian distributions separated by less than FWHM produce flat top when combined.
Model:
Perfect gaussians with uniform offsets

## Slide 34: DAC DNL is also measured
DAC DNL creates low frequency ripples in ADC histogram at predictable periodicty given by number of ADC counts produced by one DAC count.
Smooth and normalize the ADC histogram to remove this small effect.
The shift in ADC mean measures the input offset DAC’s DNL.

## Slide 35: Many topics not covered
Noise power spectrum measurement.
Banding and drift.
Fixed patterns; line start transient
Interference (moving patterns)
Timing jitter
Clipping levels; dynamic range.
DAC range
Bias and clock stability
Gain stability
Linearity
Crosstalk

## Slide 36: Additional SLIDES

## Slide 37: Vital signs – CCD in the dark
Cosmic rays and overscan edges?
Reset and clock feedthroughs normal?
Serial and parallel clocks connected?
  Inter-phase coupling
  Clock capacitance
  Resistance from one side of CCD to other
Finding CCD reset threshold.
Source follower gain and output impedance.
Source follower DC output level.
Internal AC coupler switch threshold

## Slide 38: Vital signs – CMOS mux, in the dark
Linecheck & frame check
Output offset normal?
Output offset response to load current setting biases
Mux gain (change Vreset)
Unsubstracted image with permanent reset.
  Map of source follower offsets makes recognizable pattern.
What if its not working? How to check signal continuity to detector?