## Slide 1
Compact Stirling cooling of astronomical detectors
Gert Raskin, Johan Morren, Wim Pessemier, Jeroen Vandersteen, Jesus Perez Padilla Institute of Astronomy, KU Leuven University, Belgium Mercator Telescope, La Palma, Spain
MAIA: Mercator Advanced Imager for Asteroseismology
3-Channel imager: U G R bands
3 2kx6k FT CCDs (CCD42-C0, e2v)
CCDs cooled to 165K by free-piston Stirling coolers: + + : Compact, low power, simple to use, (long lifetime) – – : Vibrations (linear cooling engine at 60 Hz)
U camera 330 – 400 nm
400 – 550 nm
G camera
575 – 700 nm R camera
Guiding camera
λ > 875 nm
Collimator
MAIA
FoV: 9.4’ x 14.1’
Mercator Telescope
Institute of Astronomy

## Slide 2
Compact Stirling cooling of astronomical detectors
Gert Raskin, Johan Morren, Wim Pessemier, Jeroen Vandersteen, Jesus Perez Padilla Institute of Astronomy, KU Leuven University, Belgium Mercator Telescope, La Palma, Spain
MAIA
FoV: 9.4’ x 14.1’
Mercator Telescope
Institute of Astronomy
Vibrations: < 10 arcsec PtP
5W lift @ 77K with 80W input power
243 mm total length 2.2 kg total mass
CryoTel MT cooler
MAIA: Mercator Advanced Imager for Asteroseismology
3-Channel imager: U G R bands
3 2kx6k FT CCDs (CCD42-C0, e2v)
CCDs cooled to 165K by free-piston Stirling coolers: + + : Compact, low power, simple to use, (long lifetime) – – : Vibrations (linear cooling engine at 60 Hz)

## Slide 3
Filter out cryo-cooler vibrations
Stirling cooler mounted on 4 leaf springs
No damping
Spring-mass system: mechanical low-pass filter (10 Hz)
60 Hz + harmonics strongly attenuated
Vibration reduction: ÷ 500 @ 60 Hz
Coolers have no longer any effect on image quality
Poster R14
Mercator Telescope
Institute of Astronomy