## Slide 1
System Tests of LSST Integrated Readout Chain Paul O’Connor for the LSST Science Raft Collaboration Brookhaven National Laboratory, Upton, NY
Laboratoire de Physique Nucléaire et des Hautes Energies, Paris
Harvard University, Cambridge, MA
University of Pennsylvania, Philadelphia, PA
SLAC National Laboratory, Menlo Park CA
Scientific Detector Workshop 2013

## Slide 2: LSST science focal plane
3.5° (64cm) diam.
3109, 0.2” (10mm) pixels
189 4K  4K CCDs
Fuly-butted mosaic (90% fill factor)
Beam geometry limits camera radial extent
2 sec readout
guider
WFS

## Slide 3: The problem
Electronics to control CCDs and process an aggregate of 1.5Gpixels/s must fit into 1/6m3 behind the focal plane
Inclusive of thermal and mechanical FPA support
Everything in cryostat

## Slide 4: The solution
A modular system of 21 “raft towers”.
Each is an autonomous, fully-testable 12K x 12K imager.
Roughly the same size and power envelope as a 16-channel SDSU controller.
Keys to compactness and low power:
  Shortest possible CCD-electronics interconnect length
  Analog CDS
  ASIC-based clock, bias, and video processing

## Slide 5: Numerology
UNIT | # Units | #CCDs | #Video Chan. | 
CAMERA | 1 | 189 | 3024 | 
RAFT | 21 | 9 | 144 | 
REB | 63 | 3 | 48 | 
ASIC | 378 | ½ | 8 | 
2010
2020
2013
2014

## Slide 6: Raft tower electronics
timing engine,
configuration
I/O
9-CCD
array
Preamp/DSI
ASIC
Clock/Bias
ASIC
DAC
DAC
heaters
RTDs
ADC
ADC
ADC
0.5MHz
OS-video
144
Clocks
72
Bias
135
Heater current
Temp sense
48
SPI
Timing
14 LVDS
Digital data
LVCMOS
50Mb/s
3.125 Gb/s mini-SAS
Clock/command
200MHz
3X RAFT ELECTRONICS BOARDS (REBs)
Power
+40A,+18A,+9A,+5D
Voltage/current/board temp readback
ASPIC
CABAC
18b
ADC
80MHz
12b
Bias, clk readback
video
clock/bias
slow control/monitoring
diagnostic
FPGA
Power
conditioning
Notes: 48 low level single ended video inputs, 144 LVCMOS single ended, 10V clocks

## Slide 7: Data acquisition, firmware, software
copper-to-optical
FPGA on REB hosts
  timing state machines (one per CCD)
  ASIC configuration and monitoring
  data encoding
  video and command interface (3.125Gbps custom protocol with virtual channels)
ATCA-based server
Each board contains a 10-G Ethernet with 9 nodes of RCE
  Dual-core (A-9) ARM processor (@ 1 GHZ)
  1 GByte of DDR3 RAM
  More than 250 DSP tiles
  Memory subsystem has > 6 GBytes/sec of I/O capacity
  Real-Time kernel or LINUX
  Full TCP/IP stack
RCE’s host all Server software
Firmware structure on FPGA

## Slide 8: Raft electronics key requirements
CCD bias | Provide programmable voltages for all CCD biases, except back depletion voltage
CCD clocking | Provide switches and programmable rail voltages/slew rate for all CCD clocks
Sequencer | Timing generator for CCD clocks, CDS switches, ADCs. Independent state machines for each CCD.
Video processing | CDS by dual slope integration for 48 video channels Pixel rate 545kpix/s Noise <9e- Crosstalk <0.2% Nonlinearity < 2%, 1 – 90ke- Max signal 175ke-
Data acquisition | 18b sampling, 48 channels @545kpix/s Serialize with ECC. No frame buffer on REB 3.125Gbps Cu link to optical TX/RX
Slow controls and monitoring | Configure ASICs Monitor board and CCD temperatures, P/S current and voltage Read serial ID chip CCD protection
Special diagnostic port | Read back CCD bias voltages 80MSa/s ADC to monitor CCD clocks via mux on CABAC
Environmental | Vacuum operation; -40C coldsink.; low outgassing; power ≤ 350mW/chan

## Slide 9: raft electronics board (REB)
420
103

## Slide 10: raft electronics board (REB)
CCD current sources
ASPIC
CABAC
buffers, ADCs
voltage regulators and references
FADC

## Slide 11: raft electronics board (REB)
STRIPE 3
STRIPE 2
STRIPE 1
Each stripe services one CCD

## Slide 12: raft electronics board (REB)
KEEPOUTS FOR HEATSINK
Thermal path

## Slide 13: raft electronics board (REB)
CCD current sources
ASPIC
CABAC
buffers, ADCs
voltage regulators and references
FADC

## Slide 14: raft electronics board (REB)
CCD current sources
ASPIC
CABAC
buffers, ADCs
voltage regulators and references
FPGA
daughterboard
Board temperature monitor
Supply current and voltage monitor
CCD
connector
HS serial I/O
Power
Test connector
STRIPE 1 POPULATED

## Slide 15: Single-CCD imaging setup
Small cryostat with engineering-grade 4Kx4K pre-production sensor
REB with 1 stripe populated
Full DAQ system 

## Slide 16: waveforms
500ns/div
Yellow: OS 100mV/div
Magenta: ASPIC out 200mV/div
Green: S3 2V/div
Blue: Reset gate 2V/div
5ms/div
5V/div
Parallel clocks
Pixel

## Slide 17: target image
8000 e-
650nm

## Slide 18: Bias and flatfield
8000 e-
650nm
Bias
650nm flat

## Slide 19: 55Fe xrays, -50V substrate bias
Single cluster (zoom)
Spectra with Ka, Kb peaks
.188’’ FWHM
Gain dispersion 1.5%
s=4mm

## Slide 20: photon transfer
NEEDS UPDATING
Full well capacity of CCD is ~140ke-; dynamic range of REB1 limited to ~ 60-80ke- withcurrent gain setup

## Slide 21: Noise, crosstalk, nonlinearity
Chart: XTALK %
 | 0.0235 | 0.0265 | 0.027 | 0.03 | 0.304 | 1 | 0.367 | 0.0824
aggressor channel
victims
* REB noise is 1.5 -2.2 e-
* REB xtalk is ≤ 0.04%
segment
480kpix/s

## Slide 22: Electronics power dissipation
Source | Peak power dissipation, W | 
 | Per REB | Per channel
CCD current source | 4.21 | .088
Video ASIC | 1.20 | .025
Clock/bias ASIC | 2.52 | .053
ADC differential buffer/driver | 0.31 | .0065
ADC | 0.36 | .0075
Power Regulators | 6.63 | .138
Control FPGA | 1.50 | .031
TOTAL Electronics Power | 16.73 | 0.349
Measured (on single-stripe REB) | 
Expose | 6.98W
Readout | 8.96W
Ave(LSST cadence) ~ 7W
30% power common
≈.35W/chan for fully-populated REB 

## Slide 23: Next steps
Pulser/
Emulator
25°
-100°
25°C
-100°
-40°
-100°C
-100°
-40°
-40°
1 REB only
1 REB warm + 1 CCD
1 REB + 1 CCD in cryostat
1 REB + 3 CCDs in cryostat
Full RTM: 9 CCDs, 3 REBs
VAC-OTM
VAC-OTM
VAC-OTM
TCM

## Slide 24: raft test cryostat

## Slide 25: mechanics and thermal

## Slide 26: CCD cooling

## Slide 27: electronics cooling

## Slide 28: cooling

## Slide 29: electronics

## Slide 30: vacuum enclosure under test

## Slide 31: proposed LSST commissioning camera (ComCam)
Raft test cryostat with 9 engineering-grade CCDs, full electronics and DAQ, LN2 cooling, shutter, corrector + filters
Installed on camera hexapod/rotator at LSST 3-mirror focus prior to full Camera readiness
30 arcmin field on-axis
Early commissioning studies

## Slide 32: LSST raft development team
I can see land…

## Slide 33: acknowledgements
This presentation has been co-authored by employees of Brookhaven Science Associates, LLC. Portions of this work are supported by the Department of Energy under contract DE-AC02-98CH10886 with Brookhaven National Laboratory. LSST project activities are supported in part by the National Science Foundation through Governing Cooperative Agreement 0809409 managed by the Association of Universities for Research in Astronomy (AURA), and the Department of Energy under contract DE-AC02-76-SFO0515 with the SLAC National Accelerator Laboratory. Additional LSST funding comes from private donations, grants to universities, and in-kind support from LSSTC Institutional Members.

## Slide 34: shameless plug
http://www.bnl.gov/cosmo2013/

## Slide 35
BACKUP

## Slide 36: Raft Tower Module
CCDs
raft baseplate
V-grooves for kinematic mount
pre-tensioning arm
FEE boards
housing (cold mass)
electronics cooling bars
Complete 144-Mpixel imager
Support sensors mechanically to meet strict coplanarity and piston tolerances
Thermal management of sensors and electronics
Protect sensor surfaces from condensable contamination
Provide bias, timing, and control signals for CCD operation
Low noise analog signal processing
Digitizing, multiplexing of pixel data
Diagnostics and slow controls
Optimized for
Compactness
Low power

## Slide 37: REB + CCD test bench components
Multimode cryostat with CCD, xray source, thermal control and monitoring
REB with adapter PCB and through-vacuum rigidflex cable

## Slide 38: Robust FEA Model: Complete RTM Assembly
Nominal thermal loads from: electronics, radiation from L3, conduction through materials, etc.
Expected thermal contact at all interfaces
Raft rim heaters at 1.0 Watts each (slightly less than ½ max)
Cryoplate at 143K
ColdPlate at 233K