## Slide 1
Theoretical Comparison
of CCD Video Processors
Dr. Simon Tulloch
University of Sheffield

## Slide 2
Reset (or Reference) pedestal
Signal pedestal
Reset event
Reset event
Charge dump
The video
processor
measures
this step size
Reset and clock-feedtrough
noise

## Slide 3
OS
OD
OS
RD
R
.
ADC
(1 sample
Per pixel)
-1
Pre-Amplifier
CCD
Inverting Amplifier
Integrator
Reset switch
Input Switch
Polarity Switch
Computer Bus
R
C
3 switches minimum
3 op-amps minimum
(in practice another switch is needed to vary gain
of pre-amp if more than one pixel speed
is required)
Correlated double sampler, Method 1:
Dual Slope Integrator (differential averager)
RL
RC=t
= width of measurement windows
(in general ~40% of pixel time)
= time between reference
and signal measurement windows

## Slide 4
OD
OS
RD
R
.
Bandwidth-
Limiting
(f3dB ~2 x fpix)
Pre-Amplifier
CCD
Hi-impedance
buffer
Clamp switch
Sample/Hold switch
Computer Bus
LP
.
RL
-
+
2 switches minimum
3 op-amps minimum
(in practice another switch is needed to vary 3dB
point of input pre-amp if more than one pixel speed
is required)
Correlated double sampler, Method 2:
Clamp and Sample
ADC
(1 sample
Per pixel)
= time between release of
Clamp and activation of Hold
H
S

## Slide 5
slope
=Gaussian white noise 15nV Hz-0.5
=flicker noise corner 150kHz
For the CCD231 the values are:

## Slide 6
At high pixel rates
we are dominated by
Gaussian white noise
At low pixel rates
we are dominated by
flicker noise

## Slide 7
OD
OS
RD
R
.
Bandwidth-
limiting
Pre-Amplifier
f3dB
CCD
Computer Bus
LP
RL
Correlated double sampler: Digital version (DCDS)
ADC
(Multiple samples
Per pixel)
Hardware simpler (at least on the analogue side).
This then allows digital synthesis of Dual Slope,
Clamp/Sample or other types of CDS.
No high-speed analogue switches required.
fADC ≥ 2.0 x f3dB

## Slide 8
Digital Synthesis : some examples
Dual Slope integrator (= Differential Averager)
Reset pedestal weights= +1
Signal pedestal weights = -1

## Slide 9
Digital Synthesis : some examples
Simplest possible DCDS with analogue prefilter
Pre-filter synthesised digitally
Clamp & Sample
Two ways to
do this.

## Slide 10
Note that if prefilter is too narrow the Point Spread Function can suffer
δ pixel
-ve signal
“leakage”
Trailing pixel
Upper 3dB too low
Lower 3dB too high
Infinite bandwidth
+ve signal
“leakage”
sig
ref
sig
ref
sig
ref
sig
ref
sig
ref
sig
ref
Note: read noise “switched off” to make effect clearer

## Slide 11
If the previous pixel waveforms are CDS processed using the Clamp&Sample technique we get:
Upper 3dB too low:
Following pixel is
below bias
Lower 3dB too high:
Following pixel is
above bias
Infinite bandwidth:
Perfect pixel delta function.
Below bias
Above bias
At bias

## Slide 12
Video bandwidth required, purely from PSF considerations:
Clamp&Sample should have
analogue bandwidth >2.6 Fpix
Dual Slope should have
analogue bandwidth >6 Fpix

## Slide 13
The bandwidth of 6x pixel rate required to give good PSF
also gives reasonable signal settling within 5% of pixel time.

## Slide 14
Various digital CDS techniques now
compared using a novel time-domain model
of the E2V CCD231 output amplifier.
Synthetic MOSFET noise waveform: “Virtual CCD oscilloscope”

## Slide 15
Build complex array
f
Real amplitudes
Imaginary amplitudes
FFT
t
Imaginary amplitudes
Real amplitudes
The real part is our MOSFET noise waveform
{200,000 point FFT
takes 6ms on a PC}

## Slide 16
Next add:
Reset noise pedestals.
Signal pedestals.
and bandwidth limit:
Add AC-coupling
Bandwidth limit the pre-amp
=CCD sensitivity mV/e-
=MOSFET Source follower gain (0.55 typ.)
( VRESET ~ 250mV for CCD231)

## Slide 17
The synthetic CCD waveforms were then analysed using the standard CDS techniques.
(floating point arithmetic with ≥ 200 samples per pixel )
Results compared the analytic models and E2V data sheet

## Slide 18
E2V data-sheet values are based on Clamp&Sample CDS
with 0.4Tpix between the two samples and a pre-filter
bandwidth=2.fpix
This analytic model suggests that Dual-
slope integration should give read noise as low as 1.3e- RMS (Controller noise not considered here)

## Slide 19
Mirrored
Gaussian
Mirrored
Exponential
Hamming
Window
(speculative)
1-Hamming
Window
(speculative)

## Slide 20
Mirrored Gaussian
and mirrored
exponential methods
give tiny advantage at
low-signal end
Differential Averager
(Dual Slope Integrator)
is the best all-round
performer.
Clamp&Sample
is the poorest
performer at all
pixel rates
Notes.
f3dB=8MHz in all cases.
Time resolution
of model=50ns.
AC coupled with
lower 3dB point at 30Hz.
Mirrored
exponential
Dual
Slope

## Slide 21
For s>>1 this method is equivalent to the
Dual-Slope method
20% clocking
time
5% settling time
5% settling time

## Slide 22
For Z=0 this method is equivalent to the
Dual-Slope method

## Slide 23
So fine tuning the Mirrored Gaussian
weights gives only a tiny improvement and then only at very-low pixel rates

## Slide 24
So fine tuning the Mirrored Exponential
weights gives only a tiny improvement and then only at very-low
pixel rates
Z=0 (equivalent to dual slope integrator)
Z ≤ 2

## Slide 25
Practical implementation of digital CDS :
- Account for more practical (i.e. lower) ADC frequencies
- Account for quantisation noise.
These are now included in the model…
Up to now the waveforms have been heavily oversampled (fADC > 200fpix) and all arithmetic has been floating point.

## Slide 26
Nyquist tells us
That fADC > 2.f3dB
Is there any advantage
to running the ADC
even faster?
[f3dB= analogue bandwidth]

## Slide 27
Small improvement
can be gained from
oversampling.
Diminishing returns
for fADC > 5.f3dB
oversampling factors

## Slide 29
Same true for mirrored
exponential method
Again, diminishing returns
for fADC > 5.f3dB

## Slide 30
Quantisation noise
Quantisation Noise
Analogue CDS processor with a single ADC sample per pixel will have a quantisation noise of 12-0.5=0.29 ADU. This adds in quadrature with the read noise.

## Slide 31
Now we quantise the synthetic CCD waveform and repeat the noise analysis
Focus in on one
pixel frequency and
two oversampling factors.
Note: the “granularity “ of the quantised
waveform is proportional to the inverse gain
of the system i.e. the e-/ADU in the image.

## Slide 32
fADC = 10. f3dB
fADC = 20. f3dB
The sample averaging will give floating point results.
We can thus get sub-ADU resolution from our ADC.
Pixel rate = 50kHz
Analogue Bandwidth (f3dB)=500kHz
CDS Method = Diff. Averager

## Slide 33
In conclusion:
DCDS reduces analogue component count and removes the need
for analogue switches.
Analogue bandwidth in a DCDS system needs to be at least 6x pixel rate from
PSF considerations.
3) ADC frequency needs to be at least 2x analogue bandwidth (as Nyquist would suggest).
A small reduction in noise can be achieved if this is increased to 5x.
Read-noise improvements are minimal if the ADC frequency is raised further.
4) Fancy DCDS weighting schemes offer insignificant improvements.
The differential averager is the best all-round performer when
implemented either digitally or with analogue circuitry.
In DCDS quantisation noise is greatly reduced which gives an effective
improvement to ADC resolution and a corresponding increase
in dynamic range.
The CCD231 should be capable of 1.3e- read noise with a zero-noise
controller (using a Differential Averager). This implies that even
with the root-2 noise hit from a differential signal chain the CCD231 should
still have an intrinsic noise floor below 2e-.

## Slide 34
If manufacturers could reduce corner frequency……………
1e- @ 50kHz

## Slide 35
Thank you!
Extended version of this presentation, together
with noise-model IDL software available at :
www.qucam.com