## Slide 1: Serial Acceptance Testing of 24 Detectors for
„Scientific Detector Workshop“, Florence, 7th - 11th Oct. 2013
Andreas Kelza
J.-L. Lizonb, R. Baconc(P.I.), M. Accardob, L. Capoanic, S. Deiriesb, C. Dupuyb, T. Fechnera, T. Jahna, A. Jarnoc, M. Loupiasc, A. Pecontalc, L. Piquerasc, R. Reissb, J. Richardc, G. Rupprechtb, M. Srivastava(a), O. Streichera, P. Weilbachera
a) Leibniz-Institut für Astrophysik Potsdam (AIP)
b) European Southern Observatory (ESO)
c) Centre de Recherche Astronomique de Lyon (CRAL)
Supported by grant no. 05A11BA2

## Slide 2: MUSE overview
MUSE, a second-generation VLT 3D-spectrograph
Field of View: 60 x 60 acrseconds
Spatial resolution: 0.2 arcseconds
Spatial elements: 60,000
Spectral range: 465nm – 930nm
24 spectrographs with detector systems
Adaptive Optics enhanced mode
The MUSE consortium:
(P.I.: R. Bacon, CRAL)

## Slide 3: MUSE detector system: AIT process
1) design @ ESO
2) head mechanical
assembly @ AIP
3) chip integration &
alignment @ ESO
4) CCD character-
isation @ ESO
5) transport to AIP
6) Acceptance test at spectrograph @ AIP
7) Integration & perfomance
verification @ CRAL.
Vacuum- and cryo-
system by ESO
8) Operating-SW by LATT,
DR-SW by AIP & CRAL
Production cycle of 24 MUSE detector heads
CCD231
4k x 4k, 15 mm

## Slide 4: Performance of 24 detectors (from VLT-TRE-MUS-1379 technical report to ESO)
Figure 1: Gain values for all quadrants and all channels. The dashed line displays the median value.
Figure 2: Maximum relative non-linearity for the 4 quadrants of each channel.
Figure 3: Measured dark current for the 24 channels. The dash line shows the specification.
Figure 6: Read-out noise measurement performed after shutting down all external electrical sources, except NGCs.
Figure 5: Readout noise for all channels. The mean value and peak-to-peak variation for the 4 quadrants are plotted. The dashed line displays the specification.
Figure 4: Measured remanence level for each channel after medium (blue) and high (red) saturation.
Gain: 1.2 ADU/e-
Non-linearity:
0.3 %
Dark current:
1.1 e-/hour
Remanence:
<0.05 e-
Readnoise @ quite
mean: 2.1 e-
Readnoise @ load

## Slide 5: Acceptance test with spectrograph
Fig 10: example of detector shift and rotation measured across the CCD.
Fig 11: spectrograph focus curves and image quality histograms in three-colours.
Fig 12: relative x and y shifts (top) and tilts (bottom) for all CCDs with allowed margins (green box).