## Slide 1: SDW 2013
Derek Ives + Gert Finger
E.S.O.

## Slide 2: SDW 2013
“Elementary, my dear Watson”
“A Detective Story”

## Slide 3: Title :- “The mystery of what killed the instrument sensitivity ?”
Characters :-
Derek “ Sherlock Holmes” Ives
“Red Herring” –
a false clue that leads towards a false conclusion
A detective story
The Victim :-
VISIR - VLT Imager and Spectrometer for mid InfraRed
“It is quite a three pipe problem, and I beg that you won't speak to me for twenty minutes”
– The Red-headed League
Dr. Gert “John Watson” Finger
The elusive Mr. “ELFN”

## Slide 4: The Prologue – improving instrument performance
Detector development – AQUARIUS (RVS)
1024 x 1024 device, 30 um square pixels
3 – 28 um wavelength coverage, AR coated
Proven Technology – MIRI heritage – Si:As IBC
Project began ~ 6-7 years ago
First light at telescope – 2012 – rapid deployment to telescope
VISIR – mid-IR Imager/Spectrometer for VLT
New Water vapour monitor
New Optics – N(10um) and Q(20um) band in one exposure (old system required 4 exposures)
VISIR at telescope

## Slide 5: Prologue - ESO Development work
Development of new mid-IR test facility at Garching (reuse of TIMMI instrument)
Filters
Polarimeter (used to change flux)
4K operation
Quick turnaround (4 hours)
Cryogenic Preamplification (50K)
NGC driven, 64 channels, 3 MHz
150 frames/second => > 300 Mbytes/s
RON ~ 200 e rms/1 Me full well
(c.f. 1000 e for last generation)

## Slide 6: The “crime scene” – Paranal, Chile
UT3
“Nothing clears up a case so much as stating it to another person” – Silver Blaze
Scene of “crime”
Crime Committed – summer 2012
Notes: 1 Jansky = non-SI unit of spectral flux density[ equivalent to 10E−26 watts per square metre per hertz

## Slide 7: The “crime” – kill the sensitivity
CRIME - Sensitivity is ~2-3 times less than with old detector
“You see, but you do not observe. The distinction is clear.” - A Scandal in Bohemia

## Slide 8: The Usual Suspects ?
The 5 suspects :-
Optics – Suspect 1, Innocent
External Electronics – Suspect 2
Gain, Linearity – Suspect 3
Very Low Q.E. – Suspect 4
Excess Noise – Suspect 5
“Eliminate all other factors, and the one which remains must be the truth” – The Sign of the Four

## Slide 9: Suspect Numbers 1 & 2 – Optics and Electronics
Bias
Signal
Read Noise Regime
Theory
Actual
Signal/Noise α 1/SQRT(Number of Reads) ?
“There is nothing more deceptive than an obvious fact” – The Boscombe Valley Mystery
  Innocent - Electronics eliminated
Read Noise / low flux regime

## Slide 10: Suspect Number 3, Gain + Linearity
INNOCENT - Linearity and Gain nominal, as per design values

## Slide 11: Suspect Number 4 - Quantum Efficiency ?
Instrument + Detector QE ~ 20 %
(6 optical elements, and calculated for 3 filters)
Seems low by factor of x 2 but not low enough
to explain loss in sensitivity
Instrument throughput (without detector) measured
Using pyro-electric detector ~ 50% (low)
Implies New detector QE > 40% (nominal)
Chart: Expected versus Measured Signal
| 16069751 | 32855639 | 50357880 | 68575840 | 87508096 | 107152450 | 127506020 | 148565320 | 170326220 | 192784130 | 215933880 | 239769930 | 264286320 | 289476690 | 315000000 | 341800000 | 369000000 | 396000000 | 425000000 | 454000000 | 484000000
Expected Signal (e) versus Measured Signal (e) | 2920000 | 5905000 | 9150000 | 12600000 | 16270000 | 20200000 | 24350000 | 28846600 | 33680000 | 37600000 | 43860000 | 49008600 | 54200000 | 59730000 | 65100000 | 70500000 | 75800000 | 81200000 | 86600000 | 91780000 | 96900000
“I never guess. It is a shocking habit — destructive to the logical faculty” – The Sign of Four
INNOCENT – Quantum Efficiency nominal, as per design values

## Slide 12: Typical operation - Evidence
Move secondary mirror
Chop faster than background
Fluctuations – cancels out.
Move telescope
Optical path not same in both
Chop positions – Nod to
remove residuals.

## Slide 13: A big clue !
Shot Noise Regime
Actual
Theory
“Any truth is better than indefinite doubt” - The Yellow Face
Read Noise Regime

## Slide 14: ELFN as a function of chop frequency
Sum of 500 differences
Difference of 500 Sums
We have a suspect – “Mr. ELFN”
0.25 Hz
0.25 Hz
OLD VISIR detector

## Slide 15: Visual Evidence
Difference of sums – low frequency noise
Sum of differences – high frequency noise
2000 frame data sequence, chopped at low and high frequencies

## Slide 16: More evidence - ELFN with detector temperature
"Data! Data! Data!“, "I can't make bricks without clay.“ - The Adventure of the Copper Beeches
No change seen with detector temperature

## Slide 17: More Evidence - ELFN with increasing flux
Increasing flux shifts frequency “knee” higher
Very important piece of evidence
'It has long been an axiom of mine that the little things are infinitely the most important.‘ – A case of Identity

## Slide 18: More Evidence – ELFN with integration time
Increasing integration time shifts
frequency “knee” lower

## Slide 19: More Evidence – De-correlation of ELFN
Increase dead-time between frames

## Slide 20: Sum up of evidence – power spectrum format
No ELFN in Dark pixels
Unit cell noise level
Different flux levels

## Slide 21: Change line of questioning - temporal correlation
Put flux on detector
Take a cube of 1024 frames
Find temporal auto-correlation for N pixels
Average N calculations and plot
Auto-correlation versus frame lag
'Having gathered these facts, Watson, I smoked several pipes over them, trying to separate those which were crucial from others which were merely incidental.' - The Crooked Man

## Slide 22: “The denouement” – real chopping with noise
Chop frequency = 2 Hz, sigma = 2.26 DN
Chop frequency = 1 Hz, sigma = 3 DN
Chop frequency = 5 Hz, sigma = 1.44 DN
Chop frequency = 10 Hz, sigma = 1.24 DN
Chop frequency = 15 Hz, sigma = 1.11 DN
“know a conjurer gets no credit when once he has explained his trick” – A study in scarlet

## Slide 23: Past Crime - Excess Low Frequency Noise (ELFN)
Stapelbroek (inventor) reports noise versus frequency issues – 1990s
'There is nothing new under the sun. It has all been done before.‘ - A Study in Scarlet
Chart: Excess Noise Ratio versus Frequency
| 193.76464999999999 | 96.882326999999989 | 48.441162999999996 | 38.752931000000011 | 24.220581999999986 | 19.376465000000035 | 9.6882326999999986 | 7.7505860999999854 | 4.8441162999999747 | 3.8752930999999977 | 1.9376465 | 1.5501172000000001 | 0.9688232700000019 | 0.77505861000000298 | 0.38752931000000107 | 0.19376465000000001
 | 1 | 1.019 | 1.0424 | 1.0724 | 1.1385 | 1.1817 | 1.56 | 1.7489 | 2.3071 | 2.6127 | 3.3717 | 3.5314 | 3.8738 | 4.0599 | 4.481 | 4.6727

## Slide 24: The prosecution’s summing up remarks
New IBC detector suffers from Excess Low Frequency Noise (ELFN).
Origin of noise is definitely the Si:As material –(Probable issues
with Blocking Layer)
Cannot be seen in DARK images, therefore CDS not helpful.
ELFN always there, seen in past generation of detectors,
but not noticed because of high chop rates.
ELFN is a function of flux and integration time.

## Slide 25: “Get out of jail cards”
Optimise detector operational parameters, minimise flux, longer integration times
(not always possible for wide band imaging, okay for spectrometer)
Increase chopping frequency of secondary mirror (> 1 metre diameter, ~ 50 kg mass
Operational limit is 5 Hz, typically used at 0.25 Hz)
New detector manufacturing run to fix problems – after modeling and detailed
understanding of problem

## Slide 26: Time off for good behaviour – “Whitening”
Image temporal de-correlation :-
Possible to de-convolve image with
temporal function to “whiten” image cubes
However simple universal whitening function
does not improve SNR
Probably due to separate ELFN for each pixel
Raw data must be stored (300 Mbytes/second, 1 hour exposures are typical)
Huge processing power – x 2 FFTs to do whitening in real time
before
after
Observed O = Intrinsic I convolved in time domain by a Kernel K
O = I ⊗ K
Kernel can be retrieved from AC then data can be deconvolved such that
I = F-1(F(O )/F(K))
(F is the Fourier transform
F-1 is the inverse Fourier transform)

## Slide 27: Drift scan/jittering
Original data cube – Noise ~ 10.9 DN
Jittered data cube – Noise ~ 3.2 DN
More Time off for good behaviour

## Slide 28
Special thanks to :-
Sherlock Holmes
(and Sir Arthur Conan Doyle)
The ESO NGC team for the great controller
and most especially Gert Finger
SDW2013