## Slide 1: Persistence Characterisation of Teledyne H2RG detectors
Simon Tulloch
European Southern Observatory

## Slide 2: Depletion Region Model
R.Smith “A theory for image persistence in HgCdTe photodiodes”, SPIE 7021-22, 2008

## Slide 3: Measurement Technique
Method: “delta” flash from LED, followed by a variable delay in which the pixel is “soaked” in the photo-charge followed by 8000s of non-destructive reads.
(dark reference window included)
Ultra-low background cryostat containing mosaic of three 5.3um H2RGs
#17308
#17306
#17310

## Slide 4: Typical result: de-trapped charge versus time since reset.
Increasing soak

## Slide 5: Can also differentiate data to yield persistence current versus time since reset.

## Slide 6: Simply comparison of persistence between devices is difficult
Below 1000s the red
chip looks best
At longer time scales its poor persistence performance becomes apparent.

## Slide 7: Long period traps the real problem, can degrade data for a whole night.

## Slide 8: Suggestion for an alternative metric for comparison of persistence behaviour
The “Tau spectra” : plot trap density as a function of the trap time constants.
Trap density electrons per pixel
Trap time constant , seconds

## Slide 9: Generation of Tau Spectra. Step 1 : represent data as a 3D pyramid
Increasing Soak
Increasing De-trap time
All traps empty
All traps full
Pyramid Edge 1 ->
information on de-trapping
time constants
Pyramid Edge 2 ->
information on trapping
time constants
˜
˜
Same data, replotted in 3D

## Slide 10: Step 2 : fit each edge with an ensemble of exponential functions.
De-trapping time
constants
Trap densities
Trapping time
constants

## Slide 11: Finally plot trap densities as a function of time constant
#17308. Worst kind of persistence: immediate charge trapping, de-trapping over tens of hours.
#17306.Trapping and de-trapping over a wide range of time constants. De-trapping is slower than trapping.
<
<
>
>

## Slide 12: Additional investigations into persistence behaviour
Electrically induced persistence (through manipulation of detector bias voltages).
Statistics of persistence charge.
Effect of cooling.

## Slide 13: Electrically Induced Persistence
Optical exposure
Electrical “exposure”

## Slide 14: Optical/Electrical comparison
optical
electrical
v
Glue void
Glue void

## Slide 15: Statistics of Persistence Charge
Persistence charge is sub-Poissonian, consistent with it being released from traps within depletion region.
System gain for persistence charge different than for photo-generated charge!
Read noise
limited regime
Poisson noise
limited regime

## Slide 16: Effect of Cooling
Cooling to 41K gives a huge improvement
with device #17308.
Note that device #17308 traps at least 2.5%
of the total photo charge during a 100s soak
when at 55K. Photometrically significant.

## Slide 17: Additional conclusions
Surprising differences between chips with near consecutive serial numbers.
“Red” device was likely dominated by a second population of traps that had very
short trap times and very long de-trap times and could only be stimulated optically.
Cooling helps reduce persistence (by a huge factor in the case of the second
trap population.

## Slide 18
END